## Audit Questions
Existing IPS Dental Audit Tool Questions/Options /Reference
Section No | Section | Question ID | Question Text | Response Options | Reference in HTM | Working Group | Reviewer Comments | Update Status | Adopted Original | New Question Content | New Reponse Options
1 | Prevention of Bloodborne virus exposure | Q1 | Does the practice have a policy and procedure/s in place for the prevention and management of blood borne virus exposure, including management of spillages, sharps and inoculation incidents in accordance with national guidance? | Yes / No | 2.6 | IPC & Built Environment | To be moved to the systems and processes section |  | Removed
1 | Prevention of Bloodborne virus exposure | Q2 | Have all staff received training in relation to the prevention and management of blood-borne virus exposure? | Yes / No | 1.22, 9.1, 9.5 | IPC & Built Environment | was unsure if 2 or 3 years for frequency- what do others think. There was a suggestion to have alligned with other training CPR, medical emergancy training |  | No | All clinical and decontamination staff recieve training in relation to prevention and management of blood-borne virus exposures every 2 years
1 | Prevention of Bloodborne virus exposure | Q3 | Have all staff at risk from sharps injuries received an occupational health check in relation to risk reduction in blood-borne virus transmission and general infection? | Yes / No | 2.6 | IPC & Built Environment |  |  | No | All staff at risk from sharps injuries have sight of risk assessments in relation to risk reduction in blood-borne virus transmission and general infection?
| Prevention of Bloodborne virus exposure |  |  |  |  | IPC & Built Environment |  |  |  | There are arrangements in place that ensures staff are dealt with appropriately in the event of a needle stick/ sharps injury, bite or a splash
1 | Prevention of Bloodborne virus exposure | Q4 | Can decontamination and clinical staff demonstrate current immunisation with the hepatitis B vaccine e.g. documentation? | Yes / No | 2.4s, 8.8 | IPC & Built Environment |  |  | No | Decontamination and clinical staff are up-to-date with their routine immunisations and additional vaccinations. Refer to Immunisation Against Infectious Disease (The Green Book) Chapter 12 Immunisation of healthcare and laboratory staff.
1 | Prevention of Bloodborne virus exposure | Q5 | Are chlorine-releasing agents available for blood /bodily fluid spillages and used as per manufacturers instructions? | Yes / No | 6.74 | IPC & Built Environment |  |  | No | Blood spillage kits and body fluid spillage kits are available for blood and body fluids which are in date.
| Prevention of Bloodborne virus exposure |  |  |  |  | IPC & Built Environment |  |  |  | Disinfectants products are bacterial and virucidal and compliant with EN14476/EN1276 standards
1 | Prevention of Bloodborne virus exposure | Q6 | Are sharps containers correctly assembled? | Yes / No |  | IPC & Built Environment |  |  | No | Sharps containers conform to the relevant standards and are correctly assembled with the lid secured.
1 | Prevention of Bloodborne virus exposure | Q7 | Are in-use sharps containers labelled with date, locality and a signature? | Yes / No |  | IPC & Built Environment |  |  | No | Sharps containers are labelled at beginning of use with date, locality and signed and dated when closed
1 | Prevention of Bloodborne virus exposure | Q8 | Are sharps containers filled beyond the indicator mark? | Yes / No |  | IPC & Built Environment |  |  | No | Sharps containers are available at point of use and positioned appropriately e.g. off the floor and away from patients when not in use
| Prevention of Bloodborne virus exposure |  |  |  |  | IPC & Built Environment |  |  |  | Sharps containers have the temporary closure mechanism in place when not in use
| Prevention of Bloodborne virus exposure |  |  |  |  | IPC & Built Environment |  |  |  | Sharps containers are not filled beyond the indicator mark
1 | Prevention of Bloodborne virus exposure | Q9 | Are sharps containers locked with the integral lock when filled to the indicator mark? |  |  | IPC & Built Environment |  |  | No | Sharps containers are locked with the integral lock when filled to the indicator mark
| Prevention of Bloodborne virus exposure |  |  |  |  | IPC & Built Environment |  |  |  | Sharps containers are locked with the integral lock when it has been used for 3 months
1 | Prevention of Bloodborne virus exposure | Q10 | Are full sharps containers stored in a secure facility away from public access? | Yes / No |  | IPC & Built Environment |  |  | No | Full sharps containers are stored in a secure facility away from public access
1 | Prevention of Bloodborne virus exposure | Q11 | Is there a readily-accessible protocol in place that ensures staff are dealt with in accordance with national guidance in the event of blood-borne virus exposure? | Yes / No | 2.6 | IPC & Built Environment |  |  | No | Needle sticks/ sharps injury, bite or splash are recorded in the accident book and reported appropriately with evidence of a full incident review
1 | Prevention of Bloodborne virus exposure | Q12 | Are inoculation injuries recorded? | Yes / No |  | IPC & Built Environment |  |  | Removed
1 | Prevention of Bloodborne virus exposure | Q13 | Are disposable needles and disposable syringes discarded as a single unit? |  |  | IPC & Built Environment |  |  | Removed
| Prevention of Bloodborne virus exposure | [8 empty] | There is evidence that allsharps are disposed of immedialty after use by person who has used it
1 | Prevention of Bloodborne virus exposure | Q14 | Are sharps containers available at the point of use and positioned safely (e.g. wall mounted)? | Yes / No |  | IPC & Built Environment |  |  | Removed
2 | Decontamination | Q1 | Does the practice have a policy or procedure that includes all appropriate aspects of decontamination within the practice e.g. cleaning, disinfection, inspection, packaging, disposal, sterilization, transport and storage of re-usable and single use instruments? | Yes / No | 2.6 | Decontamination |  |  | Yes |  | Yes/No
| Decontamination |  |  |  |  | Decontamination |  |  |  | All staff know how to access the policies and there is documentation to show that staff have read and understood the policies.
2 | Decontamination | Q2 | Has all relevant staff received training for the decontamination procedures which they are expected to perform including correct use of equipment? | Yes / No | 1.22, 2.4q, 3.7, 3.16 | Decontamination |  |  | No | Have all staff received induction and ongoing infection prevention and control (IPC) training, including decontamination procedures relevant to their role, delivered by a competent person? | Yes/No
2 | Decontamination | Q3 | Is a record kept of any instruments that cannot be reprocessed in accordance with your local decontamination policy? | Yes / No |  | Decontamination |  |  | No | Is there a documented procedure for managing instruments or items that cannot undergo standard decontamination (e.g., heat-sensitive items), in line with the manufacturer’s instructions for use, HTM 01-05 guidance, and including consideration of single-use alternatives where appropriate? | Yes/No
2 | Decontamination | Q4 | Are all wrapped, sterilized instruments dated with the use-by date? | Yes / No | 1.8, 4.24, 4.28 | Decontamination |  |  | No | Are all wrapped, sterilized instruments dated with the use-by date or Sterilisation date depending on local policy. | Yes/No
| Decontamination |  |  |  |  | Decontamination |  |  |  | Are sterilised instrument loads inspected and formally released by a designated operator, with a documented system (e.g., initials, electronic record, or log) to confirm they are fit for use?
2 | Decontamination | Q5 | Does the practice have a nominated lead responsible for infection control and decontamination? | Yes / No | 2.4c, 9.3 | Decontamination |  |  | No | Does the practice have a nominated lead with suitable and sufficient training, responsible for infection control and decontamination? | Yes/No
2 | Decontamination | Q6 | Has the registered manager a written statement of duties with specific reference to equipment validation? | Yes / No | 11.5 | Decontamination |  |  | No | Has the registered manager a written statement of duties with specific reference to equipment control including validation, servicing and maintenance?
2 | Decontamination | Q7 | Is there a procedure for transportation of instruments to and from other locations that ensures the segregation of contaminated instruments from clean/sterilised instruments? |  |  | Decontamination |  |  | No | Is there a documented procedure for the safe transportation of instruments that ensures the segregation of contaminated and clean/sterilised items during transfer within or between locations? * Transportation procedures should clearly seperate clean and contaminated instruments at all times. Procedures shoudldensure suitable containers(lidded, colour coded, leak proof). Policy should be applied to movement between treatment rooms and decontamination area or to external locations.
| Decontamination |  |  |  |  | Decontamination |  |  |  | Are cleaned and disinfected instrument transport boxes handled without gloves during movement, in accordance with infection prevention principles, to avoid cross-contamination? *Transport boxes used for moving instruments are cleaned and disinfected with a universal disinfectant wipe before being moved. *Staff understand that once disinfected, boxes can be handled without gloves. *Gloves should only be used when handling contaminated instrument boxes, not cleaned/disinfected ones. *SOPs clearly distinguish clean vs contaminated transport procedures.
2 | Decontamination | Q8 | Are all log books including testing, service, maintenance and repair records retained in the practice for at least 2 years? | All records available, Most records available, Some records available, No records available | 3.19, 4.3, 4.12, 4.14 | Decontamination |  |  | No | Are records of testing, servicing, maintenance, validation, and repairs for decontamination equipment maintained and retained in the practice for at least two years in accordance with HTM 01-05?
2 | Decontamination | Q9 | Cleaning: Are disposable instrument trays used or if re-usable trays are used are they decontaminated and sterilized after each use? | Yes / No | 2.11 | Decontamination |  |  | Yes
2 | Decontamination | Q10 | Cleaning: Are any instruments (used or unused) left on trays at the end of each session decontaminated (washed and sterilized) before further use? | Yes / No | 2.4k, 2.10 | Decontamination |  |  | No | Are all instruments, whether used or unused, that remain on trays at the end of a session decontaminated (cleaned and sterilised) before any further use? * Any instruments exposed during a clinical session must be considered contaminated and fully reprocessed before reuse. * Applies regardless of whether they were used or remained untouched.
2 | Decontamination | Q11 | Cleaning: Are instruments that are not decontaminated immediately, kept moist until they are decontaminated? | 0 hours, 1-3 hours, 3-6 hours, 6+ hours, N/A | 3.5 | Decontamination |  |  | No | If immediate decontamination is not possible, are used instruments kept moist but not immersed until cleaning and decontamination is undertaken? *Instruments should not be allowed to dry between use and cleaning *Acceptable methods:place in moist environment, use gels/sprays designed for this use *Minimise the delay where ever possible
2 | Decontamination | Q12 | Cleaning: Are Instruments inspected under illuminated magnification device for cleanliness and condition following cleaning? | Always, Almost always, Never, N/A |  | Decontamination |  |  | No | Following manual cleaning, ultrasonic cleaning, or automated washer-disinfector processing, are instruments inspected for cleanliness and condition using an illuminated magnification device? Inspection must take place after cleaning and before sterilisation. * An illuminated magnifier should be used to confirm: Instruments are visibly clean with no residual contamination. Instruments are free from damage (pitting, corrosion, wear). * Damaged instruments must be removed from service.
2 | Decontamination | Q13 | Cleaning: Are hand pieces decontaminated in-between each patient in accordance to manufacturer's instructions? | Yes / No | 2.4t | Decontamination |  |  | No | Are dental handpieces cleaned using validated processes, with correct accessories and in accordance with the manufacturer’s instructions? *Handpieces must be cleaned and sterilised after each patient use. *Decontamination processes must follow the manufacturer’s validated instructions
2 | Decontamination | Q14 | Cleaning: Are separate canisters of lubricant used for unclean, cleaned and sterilized instruments? | Yes / No | 3.56 | Decontamination |  |  | No | Are lubricant products (e.g., sprays, canisters) stored and used in a way that prevents cross-contamination between contaminated, cleaned, and sterilised instruments, in accordance with manufacturer’s instructions ?
2 | Decontamination | Q15 | Cleaning: Are those handpieces that are manually cleaned/wiped, lubricated with oil before steam sterilization in accordance with manufacturer's instructions? |  | 18.0, 3.24, 3.55, 3.56 | Decontamination |  |  | No | For handpieces that are manually cleaned in accordance with the manufacturer’s instructions (internally and externally), is lubrication with manufacturer-approved oil performed before steam sterilisation, using the correct adaptors to ensure effective penetration and performance?
2 | Decontamination | Q16 | Cleaning: Are those hand pieces decontaminated by an automated washer disinfector lubricated with oil before steam sterilization in accordance with manufacturer's instructions? |  | 3.24 | Decontamination |  |  | Removed
| Decontamination |  |  |  |  | Decontamination |  |  |  | For Handpieces processed in WD, do they carry the suitable approval symbol for automated processing.
| Decontamination |  |  |  |  | Decontamination |  |  |  | For handpieces processed in a washer-disinfector, are the correct carriers, irrigation connections, or adaptors used to ensure effective internal cleaning and decontamination, in accordance with the manufacturer’s instructions?
2 | Decontamination | Q17 | Cleaning: Are those handpieces decontaminated by an automated washer-disinfector with a specific handpiece irrigation system, lubricated with oil before steam sterilization in accordance with manufacturer's instructions? |  | 3.21, 3.24 | Decontamination |  |  | No | For handpieces decontaminated using an automated washer-disinfector, is lubrication with oil carried out before steam sterilisation in accordance with the manufacturer’s validated instructions?
2 | Decontamination | Q18 | Cleaning: Are those dental handpieces washed by a specific handpiece washer device, lubricated with oil before steam sterilization in accordance with manufacturer's instructions? |  | 3.22 | Decontamination |  |  | Removed
2 | Decontamination | Q19 | Cleaning: Are all other dental instruments washed in a washer-disinfector before steam sterilization? |  | 3.1, 3.2, 3.42, 4.3 | Decontamination |  |  | No | Are all reusable dental instruments processed in a washer-disinfector before steam sterilisation, unless validated manufacturer instructions specify otherwise, and where a washer-disinfector is not available, is an alternative validated process (e.g., ultrasonic cleaning) used, supported by documented risk assessment and SOPs
2 | Decontamination | Q20 | Cleaning: Where practices do not have a washer-disinfector, are all instruments cleaned manually before steam sterilization? |  | 2.4h, 3.33, 3.42 | Decontamination |  |  | Yes
2 | Decontamination | Q21 | Manual Cleaning: Are two sinks or two bowls in a single sink unit, used for cleaning - one for washing and a separate one for rinsing? | Yes / No | 2.4u, 5.7, 16.1 | Decontamination |  |  | Yes
| Decontamination |  |  |  |  | Decontamination |  |  |  | Are sinks and surrounding surfaces in the decontamination area maintained in a good condition, free from tiles, cracks, and clutter, and designed to be easily cleanable?
2 | Decontamination | Q22 | Manual Cleaning: Are detergents used specifically formulated for purpose of cleaning instruments? | Yes / No | 16.3, 16.3a | Decontamination |  |  | No | Manual Cleaning: Are detergents used specifically formulated for purpose of cleaning instruments and comply with the instrument manufacturer’s IFU?
2 | Decontamination | Q23 | Manual Cleaning: Is detergent used at a specified concentration according to manufacturer's guidance? | Yes / No | 16.3a | Decontamination |  |  | No | Where manual cleaning is undertaken, is the detergent used at the concentration, dilution, temperature, and as specified in the manufacturer’s instructions for use (IFU)?
2 | Decontamination | Q24 | Manual Cleaning: Is the temperature of water 45C or lower? | Yes / No | 16.3b | Decontamination |  |  | No | During manual cleaning, is the water temperature maintained at or below 45°C to prevent protein coagulation ( monitored with a water thermometer)?
2 | Decontamination | Q25 | Manual Cleaning: Where manufacturer's instructions permit are instruments fully submerged when cleaned? | Yes / No | 16.3c | Decontamination |  |  | No | During manual cleaning, are instruments fully immersed in cleaning solution where permitted by the manufacturer’s instructions for use (IFU), with exceptions such as dental handpieces processed strictly in line with their IFU?
2 | Decontamination | Q26 | Manual Cleaning: Are brushes used to clean instruments either single use or washed after each use and replaced at manufacturer's recommended interval or when damaged? | Yes / No | 16.3f | Decontamination |  |  | No | Are purpose-designed instrument cleaning brushes used during manual cleaning, and either single-use or decontaminated after each use and replaced according to the manufacturer’s instructions or when damaged?
2 | Decontamination | Q27 | Validation and Testing: Are there contractual arrangements to ensure all steam sterilizers are routinely maintained and validated in accordance with HTM requirements or with manufacturer's instructions? | Yes / No | 1.9, 11.1, 12.0 | Decontamination |  |  | No | Validation and Testing: Are there contractual arrangements to ensure all steam sterilizers are routinely maintained and validated, with completion of PVI in line with insurance requirements and manufacturer's instructions?
2 | Decontamination | Q28 | Validation and Testing: Are daily, weekly, quarterly and annual inspection, testing and maintenance records available for steam sterilizers as described in section 12 of the HTM 01-05? | All records available, Most records available, Some records available, No records available | 12.0 | Decontamination |  |  | No | Validation and Testing: Are daily, weekly, quarterly and annual inspection, testing, maintenance and PVI records available for steam sterilizers as per insurance requirements and manufacturers instructions?
2 | Decontamination | Q29 | Validation and Testing: Is the steam sterilizer removed from service following an unsatisfactory test result until the fault is rectified? | Yes / No | 11.2, 4.21 | Decontamination |  |  | No | If a steam sterilizer produces an unsatisfactory test result, is an immediate repeat test performed to confirm the failure, and if failure is confirmed, is the unit withdrawn from use until the fault is rectified and satisfactory retest results are achieved?
2 | Decontamination | Q30 | Validation and Testing: Are there arrangements to ensure all ultrasonic cleaners are maintained and validated in accordance with HTM 01-05 requirements or with manufacturer's instructions? |  | 14.0, 15.6 | Decontamination |  |  | No | Validation and Testing: Are there arrangements to ensure all ultrasonic cleaners are maintained and validated in accordance with manufacturer's instructions?
2 | Decontamination | Q31 | Validation and Testing: Are daily, weekly, quarterly and annual inspection, testing and maintenance records available for ultrasonic cleaners? | All records available, Most records available, Some records available, No records available, N/A | 14.0 | Decontamination |  |  | No | Validation and Testing: Are daily, weekly, quarterly and annual inspection, testing and maintenance records available for ultrasonic cleaners as per manufacturers instructions.
2 | Decontamination | Q32 | Validation and Testing: Are there contractual arrangements in place to ensure all automated washer-disinfectors are routinely maintained and validated in line with manufacturer's instructions? |  | 13.0 | Decontamination |  |  | No | Validation and Testing: Are there contractual arrangements in place, with a comptent person/ organisation to ensure all automated washer-disinfectors are routinely maintained and validated in line with manufacturer's instructions?
2 | Decontamination | Q33 | Validation and Testing: Are daily, weekly, quarterly and annual validation and testing results recorded for automated washer-disinfectors? | All records available, Most records available, Some records available, No records available, N/A | 13.0 | Decontamination |  |  | No | Validation and Testing: Are daily, weekly, quarterly and annual validation and testing results, as per manufacturers validation schedule recorded for automated washer-disinfectors?
2 | Decontamination | Q34 | Ultrasonic Cleaners: Are instruments placed in instrument basket or cassettes and fully immersed ensuring that all surfaces are in contact with the solution? |  |  | Decontamination |  |  | No | Ultrasonic Cleaners: Are instruments placed in instrument basket or cassettes and fully immersed ensuring that all surfaces are in contact with the solution and the bath is not overloaded. * Check instructions for max load / configuration
2 | Decontamination | Q35 | Ultrasonic Cleaners: Is the lid closed during cleaning cycles and whilst not in use to prevent contamination of the ultrasonic cleaning solution? |  | 3.30h | Decontamination |  |  | No | Ultrasonic Cleaners: Is the lid closed during cleaning cycles to prevent aerosol release?
2 | Decontamination | Q36 | Ultrasonic Cleaners: Is the water in the chamber emptied when visibly contaminated or otherwise at the end of every clinical session? |  | 3.30k | Decontamination |  |  | Yes
2 | Decontamination | Q37 | Ultrasonic Cleaners: Where instruments are manually cleaned are they rinsed after removing from the ultrasonic cleaner before sterilisation? |  | 3.30m, 3.31 | Decontamination |  |  | No | Ultrasonic Cleaners: Where instruments are processed in the ultrasonic are they rinsed after removing from the ultrasonic cleaner , with fresh potable water, and inspected before sterilisation?
2 | Decontamination | Q38 | Washer-Disinfectors: Are relevant staff aware of the instrument loading procedure i.e. spray arms are free to rotate, cannulated instruments are correctly loaded? |  | 3.17 | Decontamination |  |  | No | Washer-Disinfectors: Are relevant staff aware of the instrument loading procedure i.e. spray arms are free to rotate, handpieces are correctly loaded?
2 | Decontamination | Q39 | Washer-Disinfectors: Are cycle parameters recorded? |  | 3.19 | Decontamination |  |  | No | Washer-Disinfectors: Are cycle parameters recorded using a validated monitoring process such as datalogger/printer for each cycle?
2 | Decontamination | Q40 | Sterilizers: Is there a record made of the date, temperature and pressure achieved and satisfactory completion of each cycle? | All records available, Most records available, Some records available, No records available | 4.3, 4.14, 4.16 | Decontamination |  |  | No | Sterilizers: Is there a digital or printed record made of the date, temperature and pressure achieved and satisfactory completion of each cycle?
| Decontamination |  |  |  |  | Decontamination |  |  |  | Are Type 5 or 6 chemical indicators (e.g., TST strips) used in conjunction with steriliser digital records to provide independent assurance of sterilisation efficacy?
2 | Decontamination | Q41 | Sterilizers: Are steam sterilizers used if fault lights are displayed? | Yes / No | 4.21 | Decontamination |  |  | Yes
2 | Decontamination | Q42 | Sterilizers: Are pre-wrapped instruments placed only in vacuum type sterilizers? |  | 4.9 | Decontamination |  |  | Yes
| Decontamination |  |  |  |  | Decontamination |  |  |  | ***** suggested question Are instruments with lumens (e.g. handpieces, aspirator tips, endodontic instruments) sterilised using a validated cycle appropriate for hollow devices (typically a vacuum autoclave), in line with the manufacturer’s instructions and UK MDR requirements?
2 | Decontamination | Q43 | Sterilizers: Is distilled water, sterile water for irrigation or reverse osmosis (RO) water used in the sterilizer? | Yes / No | 4.11 | Decontamination |  |  | Yes | Sterilizers: Is distilled water, sterile water for irrigation or reverse osmosis (RO) water used in the sterilizer?
2 | Decontamination | Q44 | Sterilizers: Are opened bottles of sterile or distilled water discarded at the end of each working day? |  | 17.4 | Decontamination |  |  | No | Are opened, unused pre manufactured bottles of RO/Distilled / Sterile water discarded at the end of the day? Where practice is using reusable plastic bottles to store locally produced RO/Distilled water, is the bottle emptied and disinfected at the end of the day?
2 | Decontamination | Q45 | Sterilizers: Is the reservoir drained and left clean and dry at the end of each day? | Yes / No | 4.11 | Decontamination |  |  | Yes
2 | Decontamination | Q46 | Decontamination Environment: Is there a zoned workflow from dirty to clean? | Yes / No | 5.3, 5.6, 5.7 | Decontamination |  |  | No | Decontamination Environment: Is there a distinct division between areas and a workflow from dirty to clean?
2 | Decontamination | Q47 | Decontamination Environment: Are there separate, dedicated decontamination room/s which are restricted to those performing decontamination duties? | Yes / No | 1.8 | Decontamination |  |  | No | Decontamination Environment: Are there separate, dedicated decontamination room/s which are restricted to those performing decontamination duties?
2 | Decontamination | Q48 | Decontamination Environment: Are decontamination areas and work surfaces clean and uncluttered? | Yes / No | 5.6 | Decontamination |  |  | No | Decontamination Environment: Are decontamination areas, walls, work surfaces and cabinetry sealed, clean and uncluttered?
| Decontamination |  |  |  |  | Decontamination |  |  |  | Is the floor surface area, clean, uncluttered and sealed.
2 | Decontamination | Q49 | Decontamination Environment: Is there adequate ventilation in the clean and dirty room/s to service W/D and sterilizer? |  | 6.42 | Decontamination |  |  | No | Is the decontamination environment provided with appropriate mechanical ventilation to ensure directional airflow from clean to dirty areas, removal of excess heat/steam, and maintenance of safe and comfortable working conditions, in line with HTM 03-01 and the Workplace Regulations, with the use of uncontrolled ventilation (e.g. desk fans or open windows) prohibited?
2 | Decontamination | Q50 | Decontamination Environment: Where full mechanical ventilation is used is direction of air, flowing from the clean area to dirty? |  | 6.44, 6.45 | Decontamination |  |  | No | Decontamination Environment: Where full mechanical ventilation is used is direction of air, flowing from the clean area to dirty? * Practice should review the air exchanges per hour
2 | Decontamination | Q51 | Decontamination Environment: Are there procedures in place for the safe transfer of instruments within the practice? | Yes / No | 2.26, 2.27 | Decontamination |  |  | No | Are there documented procedures for the safe transfer of instruments within the practice, ensuring segregation of clean and contaminated items, use of dedicated transport boxes, and correct handling practices — with contaminated boxes transferred using PPE, and cleaned/disinfected boxes transferred without gloves to avoid re-contamination?
2 | Decontamination | Q52 | Decontamination Environment: Are instruments maintained in a moist condition between use and decontamination? | Yes / No | 2.12, 3.5, 3.6 | Decontamination |  |  | Removed
2 | Decontamination | Q53 | Decontamination Environment: If transport containers are in use, are they lidded, clean, leak proof and in good working order? |  | 2.25 | Decontamination |  |  | No | Decontamination Environment: Are transport containers easily identifiable for purpose, lidded, clean, leak proof, impact resistant and in good condition?
2 | Decontamination | Q54 | Decontamination Environment: Are transport containers cleaned, disinfected and dried following each use? |  | 2.27 | Decontamination |  |  | Yes | Decontamination Environment: Are transport containers cleaned, disinfected and dried following each use?
2 | Decontamination | Q55 | Decontamination Environment: Are instruments processed in non-vacuum (type N) sterilizer dried prior to packing using disposable non-linting cloth? |  | 4.23 | Decontamination |  |  | Yes
2 | Decontamination | Q56 | Decontamination Environment: Does the practice have a system in place to ensure that storage of non-wrapped instruments does not exceed: 1 day if stored in a clinical area; or 1 week if stored in a non-clinical area (i.e. clinical area not in current use, or in the clean area of a decontamination room)? | Yes / No | 4.24, 4.29 | Decontamination |  |  | Yes
2 | Decontamination | Q57 | Decontamination Environment: Does the practice have a system in place to ensure that the storage of wrapped instruments does not exceed 1 year? |  | 1.8, 4.24, 4.29 | Decontamination |  |  | Yes
2 | Decontamination | Q58 | Decontamination Environment: For each instrument is there a system in place to identify storage time, including the date by which they should be used or reprocessed? | Yes / No | 4.24, 4.26 | Decontamination |  |  | Yes
2 | Decontamination | Q59 | Decontamination Environment: Are instruments stored in a dedicated secure, dry and cool environment? | Yes / No | 4.27 | Decontamination |  |  | No | Decontamination Environment: Are instruments stored in a dedicated secure, dry and cool environment, protected from air-bourne contamination?
3 | Environmental design and cleaning | Q1 | Does the practice have a policy and procedure for cleaning and maintaining the environment? | Yes / No | 2.6, 6.54 | IPC & Built Environment | Move to systems and processes section |  | Removed
3 | Environmental design and cleaning | Q2 | Have staff undertaking cleaning duties been fully trained to undertake such duties? | Yes / No | 6.55 | IPC & Built Environment |  |  | No | All staff undertaking cleaning duties have been fully trained to undertake such duties
3 | Environmental design and cleaning | Q3 | Is the overall appearance of the clinical and decontamination environment tidy and uncluttered? | Yes / No | 5.6 | IPC & Built Environment |  |  | No | The overall appearance of the clinical and decontamination environment is tidy and uncluttered?
3 | Environmental design and cleaning | Q4 | Is the dental chair cleaned between each patient? | Yes / No | 6.46, 6.62 | IPC & Built Environment |  |  | Removed
3 | Environmental design and cleaning | Q5 | Is the dental chair free from rips or tears? | Yes / No | 6.62 | IPC & Built Environment |  |  | No | All furniture is intact without any damage and is free from rips or tears?
3 | Environmental design and cleaning | Q6 | Are all surfaces i.e. walls, floors, ceilings, fixtures and fittings and chairs free from damage and abrasion? | Yes / No | 6.38 | IPC & Built Environment |  |  | No | All surfaces such as walls, ceilings, floors are in a good state of repair with no visible signs of damage
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | All Fixtures and fittings are in a good state of repair with no visible signs of damage
3 | Environmental design and cleaning | Q7 | Are all work-surface joints intact, seamless with no visible damage? | Yes / No | 6.46, 6.47 | IPC & Built Environment |  |  | No | Are clinical surface joints are in intact with no visible signs of damage?
3 | Environmental design and cleaning | Q8 | Are all surfaces i.e. walls, floors, ceilings, fixtures and fittings and chairs free from dust and visible dirt? | Yes / No | 6.38 | IPC & Built Environment |  |  | No | All surfaces i.e. walls, floors, ceilings, fixtures and fittings are free from dust and visible dirt?
3 | Environmental design and cleaning | Q9 | Are the surfaces of accessible ventilation fittings/grills cleaned at a minimum weekly? |  | 6.64 | IPC & Built Environment |  |  | No | Ventilation fittings/grills are clean and on a documented cleaning schedule
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | Window blinds are clean and on a documented cleaning schedule
3 | Environmental design and cleaning | Q10 | Are all surfaces in clinical and decontamination areas impervious and easy to clean? | Yes / No | 6.46, 6.64 | IPC & Built Environment |  |  | No | All surfaces in clinical and decontamination areas are impervious and easy to clean
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | Mitigations are in place if surfaces are made of materials that are difficult to clean
3 | Environmental design and cleaning | Q11 | Are keyboard covers or "easy-clean" waterproof keyboards used in clinical areas? |  | 6.66 | IPC & Built Environment |  |  | No | Computer keyboard and mouse covers are "easy-clean" and waterproof in clinical areas?
3 | Environmental design and cleaning | Q12 | Are rooms where clinical practice takes place carpeted? | Yes / No | 6.46 | IPC & Built Environment |  |  | Removed
3 | Environmental design and cleaning | Q13 | Do all floor coverings in clinical and decontamination areas have coved edges that are sealed and impervious to moisture? | Yes / No | 6.47 | IPC & Built Environment |  |  | No | All floor coverings in clinical and decontamination areas have coved edges, are impervious, are sealed, and easy to clean.
3 | Environmental design and cleaning | Q14 | Are soft toys available? | Yes / No | 6.73 | IPC & Built Environment |  |  | No | All sensory/calming toys/aids used during treatment e.g. fidgets are cleaned after each use and are in a good state of repair and stored clean and dry
3 | Environmental design and cleaning | Q15 | Are free standing or ceiling mounted fans used in clinical/ decontamination areas? | Yes / No | 6.40 | IPC & Built Environment |  |  | No | The use of free standing fans and air handling units are risk assesseed before use ensuring they are used correclty and safely | Yes/No/Not Applicable
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | Fans and/or air handling units are visibly clean and on a documented cleaning schedule | Yes/No/Not Applicable
3 | Environmental design and cleaning | Q16 | Are records of cleaning maintained in accordance with HCAI Code of Practice? | Yes / No | 6.54 | IPC & Built Environment |  |  | No | Documented cleaning scehedules are maintained in accordance with HCAI Code of Practice and national cleaning standards
3 | Environmental design and cleaning | Q17 | Is cleaning equipment colour-coded, in accordance with the National Patient Safety Agency recommendations as detailed in HTM 01-05? | Yes / No | 6.53 | IPC & Built Environment |  |  | No | Cleaning equipment is colour-coded as per national cleaning standards: Red for bathrooms and toilets Blue general areas Green catering areas and kitchens Yellow known infection or isolation areas
3 | Environmental design and cleaning | Q18 | Is cleaning equipment stored in a non-clinical area? | Yes / No | 6.60 | IPC & Built Environment |  |  | No | Cleaning equipment is stored in a non-clinical area, in a locked cupboard which is labelled
3 | Environmental design and cleaning | Q19 | Where disposable single-use covers are used, are they discarded after each patient contact? |  | 6.65 | IPC & Built Environment |  |  | No | Where purpose made disposable single-use covers are used they are discarded after each patient contact, and are not used as a substitiute for regular cleaning.
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | Purpose made disposable covers are removed and surfaces are cleaned after each patient contact
3 | Environmental design and cleaning | Q20 | Are the surfaces of equipment cleaned between each patient (E.g. work surfaces, dental chairs, curing lamps, delivery units, inspection handles and lights, spittoons, external surface of aspirator and X-ray heads)? | Yes / No | 6.62 | IPC & Built Environment |  |  | No | The surfaces of equipment are cleaned between each patient (E.g. work surfaces, dental chairs, curing lamps, delivery units, inspection handles and lights, light arm, spittoons, external surface of aspirator, X-ray heads and x-ray arm, Xray control panel/buttons)
3 | Environmental design and cleaning | Q21 | Are all taps, drainage points, splashbacks, sinks, aspirators, drains, spittoons, cleaned after every session with a surfactant/ detergent? | Yes / No | 6.63 | IPC & Built Environment |  |  | No | Across all areas of the practice, taps, drainage points, splashbacks, sinks, aspirators, drains, spittoons, are cleaned after every session with a surfactant/ detergent?
3 | Environmental design and cleaning | Q22 | Are floors, cupboard doors and accessible high level surfaces and floors cleaned daily? | Yes / No | 6.63 | IPC & Built Environment | split high and low areas into seperate criteria |  | No | Low surfaces such as floors, coving and skirting boards are visibly clean and on a regular documented cleaning schedule
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment |  |  |  | High surfaces such as tops of cupboards and door frames are visibly clean and on a regular documented cleaning schedule
3 | Environmental design and cleaning | Q23 | Are floor coverings in clinical and decontamination areas impervious and easy-to-clean? | Yes / No | 6.46, 6.47, 6.49 | IPC & Built Environment |  |  | Removed
3 | Environmental design and cleaning | Q24 | Is there a designated area for the disposal of dirty water, which is outside the kitchen, clinical and decontamination areas; for example toilet, drain or slop-hopper (slop hopper is a device used for the disposal of liquid or solid waste)? | Yes / No |  | IPC & Built Environment |  |  | No | There is a designated area for the disposal of dirty water, which is outside the kitchen, clinical and decontamination areas; for example a lidded toilet, drain or slop-hopper (slop hopper is a device used for the disposal of liquid or solid waste)?
3 | Environmental design and cleaning | Q25 | Does the practice have a local policy and procedure/s for spillage in accordance with COSHH? | Yes / No | 2.4d, 2.6 | IPC & Built Environment | This is to go in the systems and processes section |  | Removed
| Environmental design and cleaning |  |  |  |  | IPC & Built Environment | added but may be in another audit section links to Q25 |  |  | Clean and dirty working areas within the clinical and decontamination rooms are clearly defined and there is a dirty to clean work flow.
4 | Hand Hygiene | Q1 | Does the practice have a local policy and procedure/s for hand hygiene? | Yes / No |  | Systems and Processes |  |  | No | Does the practice have a local or national policy and procedure/s for hand hygiene? | Yes/No
4 | Hand Hygiene | Q2 | Is hand hygiene an integral part of staff induction? | Yes / No | 6.3 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q3 | Is hand hygiene training provided periodically throughout the year? | Yes / No | 1.22, 6.3 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q4 | Is hand hygiene carried out before and after every new patient contact? | Yes / No | Appendix 1 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q5 | Is hand hygiene performed before donning and following the removal of gloves? | Yes / No |  | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q6 | Do all staff involved in any clinical and decontamination procedures have short nails that are clean and free from nail extensions and varnish? | Yes / No |  | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q7 | Do all clinical and decontamination staff remove wrist watches, wrist jewellery, rings with stones during clinical and decontamination procedures? | Yes / No | 6.9, 6.22 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q8 | Are there laminated or wipe-clean posters promoting hand hygiene on display? | Yes / No | 6.12 | Systems and Processes |  |  | No | Are there laminated, wipe-clean posters or instructions printed on soap or hand rub dispensers promoting hand hygiene on display?
4 | Hand Hygiene | Q9 | Is there a separate dedicated hand basin provided for hand hygiene in each surgery where clinical practice takes place? | Yes / No | 2.4g, 6.10 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q10 | Is there a separate dedicated hand basin available in each room where the decontamination of equipment takes place? | Yes / No | 2.4u, 5.7, 6.10 | Systems and Processes |  |  | No | Is there a separate dedicated hand basin available in each room where the decontamination of equipment takes place? If not, how is the risk mitigated?
4 | Hand Hygiene | Q11 | Are wash-hand basins free from equipment and other utility items? | Yes / No | 2.4g, 5.7 | Systems and Processes |  |  | No | Are wash-hand basins free from equipment, cleaning products, and other non-handwashing items?
4 | Hand Hygiene | Q12 | Is bar soap available at wash-hand basins? | Yes / No |  | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q13 | Are hand hygiene facilities clean and intact (check sinks taps, splash backs, soap and paper- towel dispensers)? | Yes / No | 6.11, 6.63 | Systems and Processes |  |  | No | Are hand hygiene facilities clean, free from limescale, in good condition (including sinks, taps, splashbacks, seals, and soap and paper towel dispensers), and covered by documented cleaning processes and a regular cleaning schedule?
4 | Hand Hygiene | Q14 | Are there plugs and overflows on wash-hand basins? | Yes / No | 6.10 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q15 | In the wash-hand basins does the water from the tap discharge away from the drain aperture? | Yes / No | 6.10 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q16 | Are elbow/wrist/foot operated, electronic mixers or thermostatically controlled taps available at all wash hand basins in clinical and decontamination areas? | All, Some, None | 6.10 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q17 | Are nailbrushes present at wash-hand basins? | Yes / No | Appendix 1 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q18 | Is there good quality, mild liquid soap dispensed from single-use cartridge or containers available at each wash-hand basin? | Yes / No |  | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q19 | Is skin disinfectant rub/gel available at the point of care? | Yes / No | Appendix 1 | Systems and Processes |  |  | Yes
4 | Hand Hygiene | Q20 | Are good quality disposable absorbent paper towels used at all wash-hand basins? | Yes / No |  | Systems and Processes |  |  | No | Are good quality disposable absorbent paper towels, in a dispensor, used at all wash-hand basins?
4 | Hand Hygiene | Q21 | Are hand-cream dispensers with disposable cartridges available for all clinical and decontamination staff? | Yes / No |  | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q1 | Does the practice have an infection control policy that includes procedures for the use, maintenance, service and repair of all medical devices? | Yes / No | 1.18, 2.4a, 2.6, 2.7, 3.54 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q2 | Does the practice identify an individual with nominated responsibility and authority to ensure that all staff comply with the medical device procedure? | Yes / No | 2.4c | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q3 | Has the practice carried out a risk assessment for legionella under the Health and Safety Commission's "Legionnaires' disease - the control of legionella bacteria in water systems Approved Code of Practice and Guidance" (also known as L8)? | Yes / No |  | Systems and Processes |  |  | No | Has the practice conducted a legionella risk assessment in accordance with the Health and Safety Commission’s 'Legionnaires' disease – The control of legionella bacteria in water systems: Approved Code of Practice and Guidance (L8)', and was this assessment carried out by a competent person?
5 | Management of Dental Medical Devices | Q4 | Has the Practice a written scheme for prevention of legionella contamination in water pipes and other water lines? | Yes / No | 6.75, 19.2 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q5 | Are all new re-usable instruments decontaminated prior to use? | Yes / No | 2.6, 3.4, 10.24 | Systems and Processes |  |  | No | Are all new re-usable instruments decontaminated prior to use, in accordance with the manufacturer's instructions?
5 | Management of Dental Medical Devices | Q6 | Are contaminated medical devices decontaminated prior to inspection, maintenance and repair? | Yes / No | 3.54 | Systems and Processes |  |  | No | Are contaminated medical devices decontaminated prior to inspection, maintenance, and repair, in accordance with the manufacturer's instructions?
5 | Management of Dental Medical Devices | Q7 | Are instruments sent for repair labelled to identify that they have been through the decontamination process? | Yes / No | 3.54 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q8 | Are single-use instruments re-processed? |  | 2.17 | Systems and Processes |  |  | Removed
5 | Management of Dental Medical Devices | Q9 | Are single-use endodontic files and reamers reused? | Yes / No | 2.21 | Systems and Processes |  |  | Removed
5 | Management of Dental Medical Devices | Q10 | Dental Radiography: Are intra-oral films, digital sensors and cassettes handled and stored safely in accordance with manufacturer's instructions and to reduce cross-infection? |  | 6.72 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q11 | Dental Radiography: Are film holders used in intra-oral radiography subject to sterilization after every patient use in accordance with manufacturer's instructions? |  | 6.72 | Systems and Processes |  |  | No | Dental Radiography: Are film holders used in intra-oral radiography sterilised after every patient use in accordance with the manufacturer's instructions, and subsequently stored appropriately? | provide link to HTM 0105 storage requirement
5 | Management of Dental Medical Devices | Q12 | Impression material, prosthetic and orthodontic appliances: Are impression materials, prosthetic and orthodontic appliances decontaminated in the surgery prior to despatch to laboratory in accordance with manufacturer's instructions? | Yes / No | 7.0 | Systems and Processes |  |  | No | Impression material, prosthetic and orthodontic appliances: Are impression materials, prosthetic and orthodontic appliances decontaminated in the surgery prior to despatch to the laboratory, in accordance with the manufacturer's instructions, and placed in a designated outgoing container where a disinfectant soak is used?
5 | Management of Dental Medical Devices | Q13 | Impression material, prosthetic and orthodontic appliances: Are prosthetic and orthodontic appliances decontaminated before being placed in the patient's mouth? | Yes / No | 7.1b | Systems and Processes |  |  | No | Impression material, prosthetic and orthodontic appliances: Are prosthetic and orthodontic appliances decontaminated before being placed in the patient’s mouth, in accordance with the manufacturer’s instructions, using a designated container where a disinfectant soak is used?
5 | Management of Dental Medical Devices | Q14 | Other Medical Devices: Are single-use items only used for single-treatment episode and disposed of following use? |  | 2.17 | Systems and Processes |  |  | No | Other Medical Devices: Are single-use items (including endodontic files and reamers) only used for single-treatment episode and disposed of following use?
5 | Management of Dental Medical Devices | Q15 | Other Medical Devices: Are Reusable endodontic reamers and files treated as single patient use? |  | 2.18, 2.19 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q16 | Other Medical Devices: Are difficult-to-clean instruments/devices (e.g. matrix bands, saliva ejectors, aspirator tips and three-in-one tips, etc) identified as single-use? | Yes / No | 2.17 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q17 | Dental Unit Water lines (DUWLs): Are in-line filters cleaned/replaced as per manufacturer's instructions? |  | 6.89, 6.90 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q18 | Dental Unit Water lines (DUWLs): Is there an independent bottled-water system used to dispense distilled, reverse osmosis (RO) or sterile water to supply the DUWL? |  | 6.84 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q19 | Dental Unit Water lines (DUWLs): For dental surgical procedures, involving irrigation is a separate single use sterile water source used for irrigation? | Yes / No | 6.91 | Systems and Processes |  |  | No | Dental Unit Water lines (DUWLs): For dental surgical procedures, involving irrigation is a separate single use sterile water or isotonic saline source used for irrigation? | yes, no, N/A
5 | Management of Dental Medical Devices | Q20 | Dental Unit Water lines (DUWLs): Are the DUWLs drained down at the end of every working day? |  | 6.82 | Systems and Processes |  |  | No | Dental Unit Water Lines (DUWLs): Are self-contained water bottles (bottled water systems) maintained in accordance with manufacturer’s guidance, including daily removal, flushing with distilled or RO water, and air drying where applicable, or managed through an alternative validated system that does not require daily draining? | yes, no, N/A
5 | Management of Dental Medical Devices | Q21 | Dental Unit Water lines (DUWLs): Are self-contained water bottles (bottled water system) removed, flushed with distilled or RO water and left open to the air for drying on a daily basis, and if necessary overnight, and in accordance with manufacturer's guidance? |  | 6.83 | Systems and Processes |  |  | No | Dental Unit Water Lines (DUWLs): Are self-contained water bottles (bottled water systems) maintained in accordance with manufacturer’s guidance, including daily removal, flushing with distilled or RO water, and air drying where applicable, or managed through an alternative validated system that does not require daily draining? | yes, no, N/A
5 | Management of Dental Medical Devices | Q22 | Dental Unit Water lines (DUWLs): Where bottled water systems are not used is there a physical air gap separating dental unit waterlines from mains water systems (Type A)? |  | 6.84 | Systems and Processes |  |  | No |  | yes, no, N/A
5 | Management of Dental Medical Devices | Q23 | Dental Unit Water lines (DUWLs): Are DUWLs flushed for a minimum of 2 minutes at start of each working day and for a minimum of 20-30 seconds between every patient? | Yes / No | 6.85 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q24 | Dental Unit Water lines (DUWLs): Are all DUWL and handpieces fitted with anti-retraction valves? | Yes / No | 6.87 | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q25 | Dental Unit Water lines (DUWLs): Are DUWLs either disposable or purged using manufacturer's recommended disinfectants? | Yes / No |  | Systems and Processes |  |  | Yes
5 | Management of Dental Medical Devices | Q26 | Dental Unit Water lines (DUWLs): Are DUWL filters changed according to the manufacturer's guidelines? | Yes / No | 6.89 | Systems and Processes |  |  | Removed
5 | Management of Dental Medical Devices | Q27 | Inhalation Sedation Machines [ISM]: Are ISM breathing systems (tubing, masks, nasal hood and nose pieces) used in accordance with manufacturers' or suppliers' instructions? |  |  | Systems and Processes |  |  | Removed
5 | Management of Dental Medical Devices | Q28 | Inhalation Sedation Machines [ISM]: Are ISM flowmeters used and maintained in accordance with original equipment manufacturers' or suppliers' instructions? |  |  | Systems and Processes |  |  | Removed
6 | Personal Protective Equipment | Q1 | Does the practice have a policy and procedures for the use of personal protective equipment? | Yes / No | 2.6, 6.13 | IPC & Built Environment | Needs moving to policy section |  | Removed
6 | Personal Protective Equipment | Q2 | Are staff trained in the use of personal protective equipment as part of the practice induction? | Yes / No | 6.13 | IPC & Built Environment |  |  | No | Staff are trained in the use of personal protective equipment as part of the practice induction
| Personal Protective Equipment |  |  |  |  | IPC & Built Environment |  |  |  | Personal protective equipment is included in education/training plans
6 | Personal Protective Equipment | Q3 | Are powder-free CE marked gloves used in the practice? | Yes / No | 6.20 | IPC & Built Environment |  |  | Removed
6 | Personal Protective Equipment | Q4 | Are alternatives to latex gloves available? | Yes / No | 6.19, 6.20 | IPC & Built Environment |  |  | No | Sterile and non-sterile (latex free) are available
6 | Personal Protective Equipment | Q5 | Are all single-use PPE disposed of after each episode of patient care? | Yes / No | 6.21, 6.25, 6.36c | IPC & Built Environment |  |  | No | There is evidence that PPE is changed between patients and between episodes of care
| Personal Protective Equipment |  |  |  |  | IPC & Built Environment |  |  |  | There is evidence that PPE is not worn outside of clincial and decontamination areas
6 | Personal Protective Equipment | Q6 | Is hand hygiene performed before donning and following the removal of gloves? | Yes / No |  | IPC & Built Environment |  |  | No | Hand hygiene is performed before donning PPE and after doffing PPE
6 | Personal Protective Equipment | Q7 | Are clean, heavy duty household gloves available for domestic cleaning and decontamination procedures where necessary? | Yes / No | 6.23 | IPC & Built Environment |  |  | No | Clean, heavy duty household gloves are available for domestic cleaning and decontamination procedures where necessary
6 | Personal Protective Equipment | Q8 | Are heavy-duty household gloves washed with detergent and hot water and left to dry after each use? |  | 6.23 | IPC & Built Environment |  |  | No | Heavy-duty household gloves are washed with detergent and hot water and left to dry after each use
6 | Personal Protective Equipment | Q9 | Are heavy-duty household gloves replaced weekly or more frequently if worn or torn? |  | 6.23 | IPC & Built Environment |  |  | No | Heavy-duty household gloves are replaced if they become damagaed
6 | Personal Protective Equipment | Q10 | Are disposable plastic aprons worn during all decontamination processes or clinical procedures where there is a risk that clothing/uniform may become contaminated? | Yes / No |  | IPC & Built Environment |  |  | No | The use of PPE (Disposable aprons, goggles, masks, face shield) are risk assessed and only worn worn during decontamination processes or clinical procedures where there is a risk that clothing/uniform may become contaminated
6 | Personal Protective Equipment | Q11 | Are single-use plastic aprons disposed of as clinical waste after each procedure? | Yes / No | 6.25 | IPC & Built Environment | Needs moving to Waste section, and made PPE rather than just aprons |  | Removed
6 | Personal Protective Equipment | Q12 | Are plastic aprons, goggles, masks or face shields used for any clinical and decontamination procedures where there is a danger of splashes? | Yes / No |  | IPC & Built Environment |  |  | Removed
6 | Personal Protective Equipment | Q13 | Are masks disposed of as clinical waste after each use? | Yes / No | 6.27, 6.36 | IPC & Built Environment |  |  | Removed
6 | Personal Protective Equipment | Q14 | Are all items of PPE stored in accordance with manufacturers' instructions? |  | 6.14 | IPC & Built Environment |  |  | No | All items of PPE stored in orginal packaging, away from sunlight and away from heat and liquids-
| Personal Protective Equipment |  |  |  |  | IPC & Built Environment |  |  |  | All PPE is stored in easily accessible cupboards or wall mounted dispensers close to the point of use
6 | Personal Protective Equipment | Q15 | Are uniforms worn by all staff changed at the end of each day and when visibly contaminated? | Yes / No | 6.34 | IPC & Built Environment |  |  | No | Clean, freshly laundrerd uniform are worn each day
6 | Personal Protective Equipment | Q16 | Is eye protection for staff used during decontamination procedures cleaned after each session or sooner if visibly contaminated? | Yes / No | 6.29 | IPC & Built Environment |  |  | Removed
6 | Personal Protective Equipment | Q17 | Is eye protection provided for the patient and staff decontaminated after each episode of patient care? | Yes / No | 6.29 | IPC & Built Environment |  |  | No | Eye protection is provided for patients and staff and is decontaminated each time it is used
7 | Waste | Q1 | Does the practice have a policy and procedure/s for the management and disposal of waste? | Yes / No |  | Systems and Processes |  |  | No | Does the practice have a policy and procedure(s) for the segregation, management, and disposal of waste?
7 | Waste | Q2 | Have all staff attended induction and ongoing training in the process of waste disposal? | Yes / No |  | Systems and Processes |  |  | No | Have all staff attended induction and ongoing training in the segregation and disposal of waste?
7 | Waste | Q3 | Is there evidence that the waste contractor is a registered waste carrier? | Yes / No |  | Systems and Processes |  |  | Yes
7 | Waste | Q4 | Is the Practice registered with Environment agency if generating over 500kg per annum of hazardous waste? |  |  | Systems and Processes |  |  | Removed
7 | Waste | Q5 | Are all disposable PPE disposed of as clinical waste? |  | 6.26, 6.27, 6.36, Figure 13 (07-01) | Systems and Processes |  |  | Yes
7 | Waste | Q6 | Are orange bags used for infectious Category B waste such as blooded swabs, blood- contaminated gloves and teeth without amalgam fillings? | Yes / No | Figure 13 (07-01), 5.39 (07-01), Ch10 - Dental 12 (07-01) | Systems and Processes |  |  | No | Are orange bags used for infectious waste (not containing amalgam) that arises from a patient being treated for an infection, from contact with a patient carrying a transmissible disease, or from items such as PPE that have come into contact with an infectious patient, as well as waste from a patient with a known history of infection or where a clinician has identified it as infectious?
7 | Waste | Q7 | Are yellow striped black bags used for offensive/hygiene waste such as non-infectious recognisable healthcare waste e.g. gowns, tissues, non-contaminated gloves, X-ray film, etc, which are not contaminated with saliva, blood, medicines, chemicals or amalgam? | Yes / No | Figure 13 (07-01), 5.50 (07-01), Ch10 - Dental 8 (07-01) | Systems and Processes |  |  | No | Are yellow striped black bags used for offensive/hygiene waste such as non-infectious recognisable healthcare waste e.g. gowns, tissues, gloves, X-ray film covers, etc, which are not contaminated with infected saliva, infected blood, medicines, chemicals or amalgam?
7 | Waste | Q8 | Are black/clear bags used for domestic waste including paper towels? | Yes / No | Figure 13 (07-01), 5.51 (07-01) | Systems and Processes |  |  | Yes
7 | Waste | Q9 | Are bins foot operated or sensor controlled, lidded and in good working order? | All, Some, None |  | Systems and Processes |  |  | Yes
7 | Waste | Q10 | Are local anaesthetic cartridges and other Prescription Only Medicines (POMs) disposed of in yellow containers with a yellow lid that conforms to BS 7320 (1990)/UN 3291? | Yes / No | Figure 13 (07-01), Ch10 - Dental 11 (07-01) | Systems and Processes |  |  | No | Are local anaesthetic cartridges and other non-cytotoxic Prescription Only Medicines (POMs) disposed of in yellow containers with a yellow lid that conforms to BS 7320 (1990)/UN 3291?
7 | Waste | Q11 | Are clinical waste sacks securely tied and sharps containers locked before disposal? | Yes / No |  | Systems and Processes |  |  | Yes
7 | Waste | Q12 | Are all clinical waste bags and sharps containers labelled before disposal? | Yes / No |  | Systems and Processes |  |  | Yes
7 | Waste | Q13 | Is waste awaiting collection stored in a safe and secure location away from the public within the practice premises? | Yes / No |  | Systems and Processes |  |  | No | Is waste awaiting collection stored in a safe, secure, and appropriately segregated location, whether inside or outside, in compliance with national healthcare waste guidance, ensuring it is protected from unauthorised access, weather, and public exposure?
7 | Waste | Q14 | Are all clinical waste bags fully described using the appropriate European Waste Catalogue (EWC) Codes as listed in HTM 07-01 (Safe Management of Healthcare Waste)? |  |  | Systems and Processes |  |  | Yes
7 | Waste | Q15 | Are all consignment notes for all hazardous waste retained for at least 3 years? | Yes / No |  | Systems and Processes |  |  | Yes
7 | Waste | Q16 | Has the practice been assured that a "duty of care" audit has been undertaken and recorded from producer to final disposal? | Yes / No |  | Systems and Processes |  |  | Removed
7 | Waste | Q17 | Is there evidence the practice is segregating waste in accordance with HTM 07-01? | Yes / No |  | Systems and Processes |  |  | Yes
| Waste |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | PPE is disposed of in correct waste receptical after each procedure
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The Practice has a policy for the management of blood borne virus exposure
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The Practice has a policy for the management of spillages
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The Practice has a policy for the management of sharps and inoculation incidents
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The practice has a policy and procedure for cleaning and maintaining the environment
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The practice has a local policy and procedure/s for spillage in accordance with COSHH?
| Policy |  |  |  |  | Moved from IPC & Built Environment to Systems and Processes |  |  |  | The practice has a policy for the use of personal protective equipment?